## Slide 1
PROGRESS REPORT
3D Foot-Trajectory Measurement
for Obstacle Crossing
Turning two consumer iPads into a validated ~2 mm 3D motion-capture system
Ehsan Sharafian · Biomechanics Lab · August 2026

## Slide 2
Goal & Approach
Measure foot clearance and placement during obstacle crossing in 3D — using cameras as the PRIMARY spatial measurement source, with IMUs for gait-event timing.
1
Two iPads (10th gen)
1080p @ 240 fps, locked focus/exposure, 1x lens
2
Camera calibration
ChArUco board → lens intrinsics + stereo geometry
3
Coloured foot markers
distinct colours (red / teal / magenta) tracked in 3D
Lab capture volume (floor-taped crossing zone), two-camera view

## Slide 3
Calibration & Accuracy Validation
≈ 2 mm
3D reconstruction accuracy
validated against a rigid wand with known marker distances
Lens intrinsics
both cameras stable & sub-pixel
Stereo geometry
0.74 px error · 2.04 m baseline
Wand accuracy test
2.1 mm (clean poses) across the volume
ChArUco calibration board (printed & measured)
Method: calibrate once → reuse for every session. Accuracy proven end-to-end before collecting any participant data.

## Slide 4
Result: 3D Marker Trajectory in Motion
1
First moving 3D trajectory
markers tracked & triangulated at every frame as the wand moves
2
Validation
the rigid wand's known distances stay flat through the motion — sync + tracking + triangulation all work together
3
Output
per-frame X / Y / Z for each marker, exported to CSV
Top: reconstructed distances stay on the known values (dotted). Bottom: the markers' 3D paths.

## Slide 5
Result: 3D Marker Trajectory in Motion
3
Top: reconstructed distances stay on the known values (dotted). Bottom: the markers' 3D paths.

## Slide 6
Status & Next Steps
DONE
Two-iPad system calibrated (intrinsics + stereo)
3D accuracy validated — ~2 mm
Coloured-marker detection & tracking
Camera-to-camera synchronisation
Moving 3D trajectory + CSV export
Trajectory cleaning (filter / gap-fill)
NEXT
Foot pilot: markers on a shoe, one obstacle crossing
IMU sync + gait events → locate each crossing
Compute clearance & foot-placement metrics
Automated one-command pipeline per participant
Scale: 6 obstacles × 5 crossings, many participants
The hardest technical risk — mm-accuracy from consumer cameras — is already retired.